## Slide 1
Photonic measurement of gas pressure.
Arbeitsgruppe 7.55
André Kussicke

## Slide 2: Motivation of the Quantum Pascal Project
K
s
m
Usefull in vacuum
Methods based on density measurements
kg
m
s
Status Quo:
Force per area
New theory
u(nHe−1) < 10−6
Helium
Frequency:
New SI: u(k) = 0

## Slide 3: Quantum Pascal Project: Goals
Design new optical pressure standards.
AG 7.54 – Vacuum Metrology
AG 3.33
Pressure (Pa)
Relative uncertainties
Status Quo – Established methods
Hg-Manometer
Schwien (BS)
Pressure-
gauge
Static Expansion
Continious Expansion

## Slide 4: New photonic methods
Hg-Manometer
Schwien (BS)
Pressure balance
Static Expansion
Continious Expansion
Cavity-Ring-Down-
Spectroscopy (CO2)
Refractivity measurements
(He, Ar & N2)
Pressure (Pa)
relative uncertainties
AG 7.54 – Vacuum Metrology
AG 3.33
Absorption-spectroscopy
with Herriott-Cells (CO2)
Magneto-
optic
traps

## Slide 5: Measurement principle
Fabry-Perot-Resonator
Measurement of the refractivity by using the frequency difference.
Calculation of the partial pressure using the ideal or real gas law.
T
Laser 2
Laser 1
Lock-In-technique and
frequency measurement
Double-resonator
Thermostatisation
Lock-In-Technique feedback loop
Pure gases
Differential pressure measurement
Gases: He, Ne, Ar
p

## Slide 6: Measurement principle
T
Laser 2
Laser 1
Lock-In-technique and
frequency measurement
Double-resonator
Thermostatisation
Lock-In-Technique feedback loop
Pure gases
Differential pressure measurement
Gases: He, Ne, Ar
p
Temperature
u(T) < 300 μK
Gas purity
7.0
Length
u(L) < 3*10-11 m
Frequency
u(f) < 15 kHz

## Slide 7: New refractometer at the PTB

## Slide 8: Pressure induced deformation

## Slide 9: Pressure induced deformation
Pressure induced cavity length change: 6.389 pm per Pa

## Slide 10
Thermodynamic effects
pV work heats up the gas and cavity walls.
The temperature change (ΔT) is maximal at 5mK (for t=1s).
After 40s ΔT is less than 0.5mK and does not negatively influence measurements.

## Slide 11: Frequenzy stabilisation
Opto-elektronic frequency stabilisation:
  Lock-in-technique
  Comparison with stabilized HeNe-Laser
  Pound-Drever-Hall-stabilisation
L1 Tuneable diode Laser
L2 Tuneable diode Laser
Detector
Detector
Lock-in-control
Spectrum Analysator
Frequency-Counter
Lock-in-control
L-2
L-2
L-2
L3 Stabilized HeNe-Laser
Beat Detector
Spectrum Analysator
Frequency-Counter
Fast Detector
Resonator
Resonator
L-2

## Slide 12: Uncertainty of refractivity measurements
Frequency shift in GHz
Pressure in kPa
Frequency shift over pressure
„DKM“ (Pressure balance)
Δ(n-1)He = 3.8*10e-4
Pressure measured with RSG‘s
Laser frequency shift obtained from beat measurement with reference resonator.
Comparison with pressure generated by pressure balance (DKM) at 30 kPa.

## Slide 13
Physikalisch-Technische Bundesanstalt
Braunschweig und Berlin
Abbe Straße 2-12
10587 Berlin
André Kussicke
E-Mail: Andre.Kussicke@PTB.de
www.ptb.de

## Slide 14: Virialkoeffizienten
Karl Jousten, et al. Perspectives for a new realization of the pascal by optical methods. Metrologia 54 S146 (2017)